**ALARIS DATASET REVISION**

**11/21/2018**

**CHANGES:**

**Drug Change \_\_\_\_\_\_\_\_ \_\_\_\_\_\_\_\_\_\_\_\_\_ Profiles\_\_\_\_\_\_\_\_\_\_**

Heparin Increase hard max to 36 units/kg/hr CC, IC, MS, IPONC, OB

Heparin Update EKOS sub-type by adding non-weight CC option if patient’s cut-off weight is reached

Heparin Add Impella sub-type CC

Heparin Non-weight based sub-type CC, IC, MS, IPONC, OB

Heparin Remove Centri-Mag sub-type CC

Magnesium infusion (TTM) Add option for Targeted Temperature CC Management (TTM) bolus and cont. infusion

Alteplase-EKOS and Add 13 mg in 250 ml option CC  
 Radiology sub-types

Drotrecogin (Xigris) Remove from library CC

Levofloxacin Add to Infusion Center’s profile INF

Ceftazidime Add to Infusion Center’s profile INF

Dipyridamole Add option for syringe module NUC

Sincalide Add option for syringe module NUC

Dobutamine Add option for syringe module NUC

Maintenance IVF Add to Nuclear Medicine’s profile NUC

CC = Critical Care

IC = Intermediate Care

INF = Infusion Center

MS = Medical/Surgical

OB = Obstetrics

IPONC = Inpatient Oncology

OPONC = Outpatient Oncology

RHU = Rheumatology

NUC= Nuclear Medicine

SCN = Special Care Nursery

PED1 = Pediatric 5 Kg to 19.9 Kg

PED2 = Pediatric 20 Kg to 39.9 Kg

PED3 = Pediatric greater than 40 Kg