ALARIS DATASET REVISION

1/30/2013

CHANGES:

Drug			Change	___			_____			Profiles

Increase Cardizem soft max to 20 mg/hr						CC, IC

Add Insulin bolus from bag on all adult profiles					CC, IC, MS, OB, ONC

Add Insulin drip option to Oncology profile						ONC

Change hard minimum concentration of KCL to 0.08 mEq/mL			CC, IC, MS, OB, ONC

Increase duration of KCL infusions by 30 minutes to 330 min			CC, IC, MS, OB, ONC

Add Precedex 200 mcg in 50 mls as option						CC

Add Precedex bolus of 1 mcg/kg over 10 minutes					CC

Increase soft max of PRBC, FFP and whole blood to 999 ml/hr			CC

Add additional selection under Heparin for CentriMag				CC

Increase bolus administration rates of Meds run on syringe module in anesthesia.	CC

Remove soft limit alerts on meds run on syringe module in anesthesia		CC

(Atracurium, Fentanuyl, Vecuronium, Remifentanil, Sufentanyl, Propofol, Cisatracurium)

CC = Critical Care

IC = Intermediate Care

MS = Medical/Surgical

OB = Obstetrics

ONC = Oncology